## Slide 1: X-ray Detection
Antonino Miceli (amiceli@aps.anl.gov)
June 14, 2015
NX School

## Slide 2: Outline
Counting vs. integrating
Indirect versus direct detection
  Scintillation Counters
  Area detectors using scintillators
    Large area for diffraction (low spatial resolution, ~ 100 mm)
    Small area for imaging (high spatial resolution, ~ 1 mm)
  Ion Chambers
  Pixel array detectors (e.g., Pilatus)
  Energy resolving detectors (i.e., spectroscopic detectors)
    Measuring the energy of photons
    Silicon diodes
    Superconducting detectors

## Slide 3: How do you detect x-rays?
Need to convert to something that you can measure
  Electrons… Q = CV
    Indirectly (x-rays  optical photons  electrons)
      Scintillators + Optics + photomultiplier/CCDs
    Directly (x-rays  electrons)
      Ion Chambers, Pixel Array detectors (e.g., Pilatus)
  Temperature
    DT = Eg / (Heat Capacity)
      Superconducting calorimeters

## Slide 4: Counting versus Integrating
Analog Output
Pulse Height ~ Energy
Integrating
Digitizer
Dark Current
Counting
threshold
?
?

## Slide 5: Counting versus Integrating
Counting
  Single photon counting
    Scintillator counting detectors (e.g., Cyberstar)
    Pilatus (counting pixel array detectors)
    Energy-resolving Detectors (Silicon or Germanium diode detectors)
  Deadtime limitations!!!
  Dark current rejected with a sufficiently high threshold.
Integrating
  Signal accumulates
  CCDs, Ion chamber
  No deadtime limitations
  Read noise and dark current are issues to consider

## Slide 6: Shaping time – counting detectors
Response time of detector
  In most detectors, the user can change this via software.
Gain is usually associated with longer shaping time.
Longer shaping time improved the energy resolution
  But reduced the total count rate throughput.
  Energy resolution is important for certain techniques (e.g., XANES, XAFS)
Analog response of a counting detector to the same x-ray photon energy

## Slide 7: Deadtime limitations for counting detectors
Analog pulses
Discriminator output (Digital)

## Slide 8: Deadtime
As you increase the input count rate (ICR), does the output count rate (OCR) follow linearly?
  The longer the shaping time, the lower the ICR before deviating from linearity.
When to worry?
  Rate > 1 / (2 x t)
Input count rate (ICR)
Output count rate (OCR)

## Slide 9: Deadtime for synchrotron (pulsed source)
Depends on the fill pattern and speed of the detector
Counting detectors can detect at most one photon per bunch
D.A. Walko, D.A. Arms, E.C. Landahl, J. Synchro. Rad. 15 (2008) 612; http://dx.doi.org/10.1107/S0909049508022358
NaI
YAP
APD
Fast detector (shorter shaping time)

## Slide 10: What fill pattern pattern will you be using?
Hybrid singlet and 324 bunch mode each 2 weeks a run.
Hybrid singlet useful for certain timing experiments.
  Not great for high count rate experiments (e.g., XANES, XAFS, spectro-microscopy)

## Slide 11: Indirectly (x-rays  optical photons  electrons) Scintillation Counters
One pixel
  “Point detector”
Still workhorse for high resolution diffraction experiments (plus a pair of slits)
NaI(Tl) is the most common scintillator and gives a energy resolution (ΔE/E) of about 35% - 40%. Organic (plastic) scintillators are used for higher speed applications but energy resolution is sacrificed.

## Slide 12: Indirectly (x-rays  optical photons  electrons) Charge Coupled Devices (CCDs)
Optical detectors are everywhere in our lives… camera phones, etc.
CCDs are integrating detectors. No dead-time issues, but read noise and dark current

## Slide 13: Indirectly (x-rays  optical photons  electrons) Charge Coupled Devices (CCDs) + x-ray scintillators
Thickness of scintillator dictates the usable energy range

## Slide 14: Indirectly (x-rays  optical photons  electrons) Charge Coupled Devices (CCDs) + x-ray scintillators
With demagnification for large area detectors
  Diffraction ( < 30 keV)
x-rays
sample
Fiber Optic Taper (Optical photons)
(1 – 3 De-Magnification)
CCD
Scintillator (e.g., Gd2O2S)
Spatial resolution ~ 100 mm
No deadtime correction
Calibrations
  Dark Subtraction
  Spatial Distortion
  Spatial gain variations

## Slide 15: Indirectly (x-rays  optical photons  electrons) Amorphous Silicon Flat Panel + x-ray scintillators
Used at higher energies ( > 50 keV)
Thin film transistor (TFT) technology (a-Si photo-sensors) allows large area detectors
  Cheaper than CCDs, but higher noise and less dynamic range!
  Typically faster than CCDs.
~ 120 cm
Sector 1

## Slide 16: Scintillator Thickness and Optimal Energy Range
Simple Energy-Dependence Model for Indirect Detection Area Detectors
Signal Size
marCCD/165
500 mm CsI
40 60 80 100 120 140
Energy (keV)
Number of X-ray Photons for S/N ratio = 1
8
7
6
5
4
3
2
GE a-Si Flat Panel
40 mm Gd2O2S
10
2
10 30 50 70 90
Energy (keV)

## Slide 17: Indirectly (x-rays  optical photons  electrons) Charge Coupled Devices (CCDs) + x-ray scintillators
Magnification  Microscopy
  mm-scale spatial resolution with x-rays

## Slide 18: Directly (x-rays  electrons) Ion Chambers
Integrating detectors… ion current ~ x-ray flux
Used to monitor beam intensity
Used to normalize data to the beam intensity (“I0”)
Also used for transmission XAS measurements.

## Slide 19: Directly (x-rays  electrons) Pixel Array Detectors (e.g., Pilatus)
Dectris Ltd

## Slide 20: Directly (x-rays  electrons) Pixel Array Detectors (e.g., Pilatus)
Pilatus is a digital PAD (photon counting)
CMOS readout chip (i.e., Application Specific Integrated circuit, ASIC)
PSI/SLS Detector Group
AGIPD pixel

## Slide 21: Directly (x-rays  electrons) Integrating Pixel Array Detectors
You can design the CMOS readout in anyway you like.
  e.g., with an integrating front end.
CSPAD at LCLS
Gruner et al.,

## Slide 22: Directly (x-rays  electrons) Pixel Array Detectors (e.g., Pilatus)
Each pixel is a single photon counting detectors!
Thus has count rate limitations
487 x 195 pixels (172 mm)
8.3 cm x 3.3 cm Area
Count Rate ~ 1 MHz/pixel
20-bit counter/pixel
5ms readout (Frame Rate = 200 Hz)
320 micron thick Silicon sensor
Gateable & electronic shutter
Lower Level Discriminator only
Brönnimann et al. @ PSI in Switzerland (Dectris)

## Slide 23: A variety of Pixel Array detectors
Quantum efficiency of pixel detectors is x-ray absorption efficiency (i.e., photoelectric)
  Desire thick and high-Z sensor materials
Silicon (Z = 14) sensors  Pilatus
  Currently maximum thickness ~ 1 mm
CdTe (Zeff = 50)sensors  Pixirad
  CdTe thickness 0.5 to 1 mm
Pixirad CdTe Pixel Detector

## Slide 24: Directly (x-rays  electrons) Pixel Array Detectors – Pilatus Threshold
Where to set the threshold?
Is there an “optimal” threshold?
Threshold

## Slide 25: Pilatus Threshold – pixel charge sharing
Pixels

## Slide 26: Pilatus Threshold – pixel charge sharing
If threshold is too high, then you under count events (effectively a small pixel)
If threshold is too low, then you double count events
“Optimal” threshold is 50% of beam energy
  Unless you need to reject fluorescent background.
B. Schmitt et al

## Slide 27: Spatial Resolution: Indirect vs. Direct
Indirect detection area detectors resolution is limited by the thickness of the scintillator because of optical blurring
Direct detection detectors have single-pixel resolution
Tungsten Knife Edge
GE a-Si Flat Panel (CsI)
Pilatus Silicon

## Slide 28: Energy Resolving Detectors (aka Energy Dispersive Detectors) (aka Spectroscopic Detectors) (aka XRF detector)
x-ray beam
Copper
XRF detector
Cu Ka
Cu Kb
Spectrum

## Slide 29: Fluorescence (XRF) Measures…
Abundance (ppm level) and spatial correlations of heavy elements
Elemental Compositions of Comet 81P/Wild 2 Samples Collected by Stardust (Flynn et al 2006) Solid-phases and desorption processes of arsenic within Bangladesh sediments (Polizzotto et al 2006)
Sea Coral
Matt Newville, 13-id
A link between copper and dental caries in human teeth identified by X-ray fluorescence elemental mapping (Harris et al 2008) Levels of Zinc, Selenium, Calcium, and Iron in Benign Breast Tissue and Risk of Subsequent Breast Cancer (Cui et al 2007)
XRF Image of a cell
Barry Lai, 2-id

## Slide 30: Spectroscopic Detectors
x-ray
X-Ray Energy ~ # of e-h pair
( 3.67 eV are need to produce 1 e-h pair for Silicon!!)
( Silicon or Germanium )

## Slide 31: In reality, we use Silicon Drift Diodes …
Noise scale like the area of the anode!
  Thus make the anode small

## Slide 32: Spectroscopic Detectors – Signal Chain
3 ms
1 V
Pulse Height
Digitizer
Pulse Height Analyzer
Multi-Channel Analyzer
Histogram
x-ray
diode
Preamp
Shaping Amp
Energy ~ P.H. ~ channel #

## Slide 33: Spectroscopic Detectors – Pulses to Histograms
3 ms
1 V
Pulse Height
Fe-55 Source on a scope.
MCA = histogram (e.g., 2048 channels)
SCA = Single Channels (i.e., ROIs)
Mn Kb
Mn Ka

## Slide 34: 4-element Silicon Drift Diode
4-element SDD (SII Nano Inc)
Best Energy Resolution ~ 150eV
Peak-to-Background Important!!!
  Usually signal is buried here!
  Recombination, incomplete charge capture, etc.

## Slide 35: Trade-off between count rate and energy resolution!!!
Shorter shaping time (length of pulse) means more count rate, but less energy resolution.
  Depends on your experiment.

## Slide 36: Dynamic Range
We typically have weak signals next to strong signals in the same image.
For counting detectors, dynamic range is set by how big your counter is (e.g., Pilatus has a 20-bit counter in each pixel)
For integrating detectors (e.g., CCD, a-Si flat panels), the dynamic range is set by how much charge each pixel can storage and the noise level.

Input Flux
Detector Signal
Level
DR = S/
S
Input Flux
Detector Signal
Level
DR = S
S
1
Gruner et al.,

## Slide 37
Beyond Silicon … Superconducting sensors
1/e response time = 100 μs - 1 ms
Kent Irvin et al, NIST-Boulder

## Slide 38
Transition Edge Sensors
High resolution, low count rates  Need to make arrays

## Slide 39
Microwave Kinetic Inductance Detectors
‘Microwave’ refers to the readout frequency!

## Slide 40
Cooper Pair
Energy Gap
Silicon – 1.10000 eV
Aluminum – 0.00018 eV
Energy resolution:
Why use Low Temperature Superconductors?

## Slide 41
Microwave Kinetic Inductance Detectors
Excess quasiparticles or DT generated by x-ray causes an inductance increase (i.e., “kinetic inductance”)
  Measure inductance change in a LC resonating circuit
Multiplexing: Lithographically vary geometric inductance/resonant frequency…
DLs
DRs
Observables….

## Slide 42
Cryogenic Detector R&D at APS
The goal is energy resolution < 5eV with good count rate capabilities (> 100kcps)
Three Main Aspects:
  Device Fabrication
    Completely in-house with dedicated deposition chamber
  Cryogenics and Device Characterization
    Turnkey 100 mK cryostat (cryogen-free)
  Readout electronics
    Multi-pixel implementation in progress

## Slide 43
Anatomy of a thermal MKID (i.e., calorimeter)
Microcalorimeter
Superconducting
Resonator
+
300 mm
Capacitor
0.5 mm thick SiN
Absorber
Inductor
Feedline
Empty
Space
0.5 x 300 x 300 mm Tantalum Absorber
100 nm WSi2 resonator

## Slide 44
Microfabration Fabrication Process
Six Layer Mask Layout (2” wafer)
0.5 mm SiN + 300 mm Silicon wafer
Resonator deposition (@ APS)
Resonator Lithography (MA-6, CNM)
Resonator Etch (Oxford RIE, CNM)
Resist strip (1165 remover, CNM)
Absorber Lithography (MA-6, CNM)
Absorber deposition (@ APS, CNM)
Absorber liftoff (1165 remover, CNM)
SiN bridge lithography(MA-6, CNM)
Backside SiN membrane lithography (MA-6, CNM)
Backside SiN etch (March etcher, CNM)
Bulk Si etch (KOH, CNM)
Backside protective Al depositions (@ APS)
SiN bridge etch (March etcher, CNM)
Al wet etch (CNM)
Resist strip (1165 remover, CNM)

## Slide 45: Conclusions
Take a moment to analyze what kind of detector you are using!
  Counting or Integrating?
    Counting: Deadtime limitations (what’s the fill pattern during my experiment?)
    Integrating: Dark Subtraction?
  Pilatus detector (counting pixel array detectors)
    What threshold should use?
  Energy resolving detectors?
    What shaping time to use?
      Speed versus resolution
  Interested in detector physics? Come talk to me!
    Looking for some young minds to develop new detectors!