## Slide 1: Evolution of the Detector Module Design for CORELLI at SNS

## Slide 2: Outline of Talk
Introduction to Corelli and Corelli Detectors
Why the Need for a Different Detector Design
Design Challenges
Working Solutions
Final Comments

## Slide 3: What is CORELLI? aka Elastic Diffuse Scattering Spectrometer (managed as part of the SING II project)
Corelli is a statistical chopper spectrometer designed and optimized to probe local disorder through diffuse scattering from single crystals
It combines the high efficiency of white beam Laue diffraction with energy discrimination by modulating the beam with a statistical ‘correlation’ chopper
Corelli will be the first instrument in the world dedicated to the study of diffuse scattering with energy discrimination

## Slide 4: CORELLI is located on BL9 at SNS
Corelli BL9
Proton Beam Direction
Bridge to CLO

## Slide 5: Instrument Layout
20 m
2.5 m
Moderator C/L
Sample C/L
Detector C/L
Choppers
Sample Scattering Vessel

## Slide 6: Detector Array
3 rows of He-3 Filled Linear Position Sensitive Detectors (GE Reuter-Stokes)
1.27 cm diameter x 83.9 cm active length, Spatial Resolution ≈ 1 cm x 1 cm
16 tubes per module, 30 modules per row for a total of 1440 tubes
Vertical Range - 28.5o to + 28.5o
Horizontal Range - 23o to + 152o
Initial phase to consist of 38 modules, mostly located along middle row
Sample Scattering Vessel
Detector Array Frame

## Slide 7: No Isolation Valve for Corelli …
No room to fit valve to isolate sample vacuum space from detector vacuum space
Thus, we need to vent and pump entire vacuum vessel with each sample change

## Slide 8: ... Resulting in a Few Design Challenges Compared with the ‘ARCS’ detector
1. To reduce pump down time, detector electronics need to be placed inside compartment which is separate from the vacuum space
ARCS design open to vacuum
2. 15 Watts per module
twice that of ARCS 8-pack design
3. Electronic circuit board designs (preamplifiers in particular) require temperatures to be stable to within 1-2 oC for reliable performance
same as for ARCS but harder
to achieve for Corelli due to
periodic vent/vacuum cycles
ARCS 8-pack detector module 2.54 cm dia. x 100 cm long tubes

## Slide 9: Design Challenge #1 Met
Electronics Tray Fits in Compartment Behind Detector Tubes
Detector Electronics Tray Assembly

## Slide 10: Design Challenge #2 (15 Watts of Power) Two Options Considered:
Active Cooling
(forced air flow through electronics compartment)
Passive Cooling via heat sinks and black surfaces
Cooling Lines
Painted Black Inside and Out

## Slide 11: Thermal Imaging Identified ‘Hot Spots’ for Placement of Heat Sinks

## Slide 12: Custom Vacuum Chamber Built to Simulate Vent/Vacuum Cycles in Lab
Atm to <1E-4 Torr in 30 min
Cooling lines connected to house air (and flow controllers)
Circuit boards have built in thermal sensors
Also used TC gauges to measure surface temperatures
Power and signal cables to quantify stability of detector electronics

## Slide 13: Design Challenge #2 Met
We targeted < 60 oC to be consistent with ARCS design
Able to adequately cool with reasonable air flow (0.5 CFM)
Also able to passively cool with black surfaces, modest repositioning of boards, and application of heat sinks
Chart: Time in hours
| 0 | 0.5 | 3.0833333333333401 | 3.5 | 4 | 4.666666666666667 | 5.4166666666666714 | 6.4166666666666714 | 6.5 | 6.8333333333333695 | 22 | 23 | 31.25 | 31.5 | 45.25 | 75.75 | 93.25
 | 30.4 | 35 | 41.3 | 42 | 42.8 | 44 | 45.3 | 46.8 | 46.8 | 46.8 | 54.3 | 54.3 | 56 | 56 | 56.5 | 57 | 57.3
Chart: Time in hours
| 2.0833333333333401 | 4.25 | 11 | 27.166666666666668 | 76.916666666666927 | 104 | 120.25
 | 47 | 47.3 | 49.3 | 52.5 | 54.3 | 54.3 | 54.3

## Slide 14: Design Challenge #3 (temperature instability of preamps)
Temperature stability during 2 hour vent cycle was not achieved
Instrument team needs to be able to take data once acceptable vacuum is reached (est. 2-3 hours)
Time constant to return to equilibrium temperatures > 24 hours
Chart: Time in Hours
| 0 | 0.16666666666666666 | 0.33333333333333331 | 1 | 2.0833333333333401 | 2.8333333333333335 | 4.25 | 8
 | 53.3 | 53 | 52.8 | 50.3 | 47 | 46.3 | 47.3 | 49.3
Repump Begins
Vent Begins

## Slide 15: Effect of Temperature Drift on Calculated Neutron Position
Temperature drift of preamps is known to have detrimental effect on calculated event position
Effect much more pronounced near tube ends than center
Chart: Temperature (oC)
| 46.3 | 47.3 | 47.8 | 48.5 | 49.8 | 56 | 48 | 48.5 | 49.3 | 49.8 | 50.5 | 52 | 58
tube center | 0 | 0.17 | 1.15 | 1.4 | 2.03 | 3.72
tube end | 0 | 4.77 | 9.5 | 14.04 | 17.41 | 26.48 | 61.72

## Slide 16: Design Challenge #3 Met with Electronic Solution
Preamplifier modification provides solution to peak shift problem
capacitive decoupling of
1st stage op amp leads to
exceptional peak stability
with temperature change
Chart: Temperature (oC)
| 46 | 46.5 | 47.5 | 48.5 | 49.3 | 56 | 47.8 | 48.8 | 49.5 | 50 | 50.5 | 51.8 | 57.8
tube center | 0 | 0.03 | -0.08 | -0.06 | 0.19 | 0.31
tube end | 0 | -0.02 | -0.16 | 0.15 | 0.04 | -0.09 | 0.11

## Slide 17: Performance Stable Over Temperature Range Expected During Sample Change-Out
units
RED (mm)
GREEN (mm)
BLUE (oC)
Chart: Time After 2 Hour Vent to Atmosphere (hours)
| 0 | 1 | 2.25 | 3.75 | 4.75 | 5.75 | 6.75 | 7.75
change in temperature (oC) | 0 | -3.3 | -6.8 | -6 | -5.5 | -5 | -4.5 | -4
change in FWHM (mm) | 0 | 0.23571 | 0.11714 | 0.47429 | 0.43857 | 0.19571 | 0.30857 | 0.4
change in peak position (mm) | 0 | 0.027143 | -0.21 | -0.16143 | -0.19 | -0.17857 | -0.03 | -0.10429
Change in FWHM (mm)
Vent Begins
Repump Begins

## Slide 18: Redesign of Preamplifier Boards
Incorporated decoupling capacitor modification
Improved protection circuitry to buffer against HV breakdown (Corona)
Plug-in connectors to facilitate tray removal for maintenance

## Slide 19: Upon further testing, vacuum group recommended removing all cabling from vacuum space
Rear view of detector module showing NW25 bellows

## Slide 20: Preparing for Installation
The first phase of 38 Corelli detector modules are currently undergoing assembly and testing, on track for an early CD-4 finish date in Feb 2014

## Slide 21: Detector Development Team – Key Players
George Rennich
  Lead Engineer
Bill Turner
  Designer
Mark Wendel
  Thermal Analysis
Kevin Berry, Justin Beal
  Detector Development
Vlad Sedov
  Electronics Development
Feng Ye
  Instrument Scientist