## Slide 1
Neutron Optics
&
Instrumentation
Bill Hamilton
Neutron Instrument & Source Division
Oak Ridge National Laboratory
Sunday 22 June 2014
National School on Neutron and X-ray Scattering
14-28 June 2014

## Slide 2: Chadwick identifies the neutron - 1932
1935 Nobel Prize in Physics "for the discovery of the neutron”
Notes: The first neutron “scattering” instrument – or at least the first neutron scattering experiment to know that it was one …Chadwick’s problem and solutions are very much ours …
In similar experiments a few years earlier Irene Curie & Frederic Joliot and Bothe & Becker a few years earlier had misidentified the neutral particle in question as photon – well gamma ray. Ettore Majorana correctly deduced as Chadwick would later that the Joliot-Curie particle had to be a massive neutral particle coming out of the nucleus otherwise the knocking on of the protons would violate the conservation of energy and momentum (also it’s a huge gamma anyway, so why didn’t the ionization chamber see it?). Thus being the neutron, something of this sort having been expected for a while - Rutherford after all had suggested in 1920 that something of this sort was needed to stabilize all the positive charges he had found concentrated by protons in atomic nuclei and the name was suggested at that time - he was Chadwick’s boss at the time. Despite this being something of a big deal and despite encouragement (cajoling?) from Fermi, Majorana didn’t bother to publish since the result was in his opinion “banal”. Chadwick did a little more than this basic experiment and by mixing targets was able to estimate the mass of the neutral particle as 1.0067 proton masses and for this got the Physics Nobel in 1935. The J-C’s btw got the Chemistry Prize for artificially induced radioactivity which is what is going on in the Be. Majorana went missing during or soon after a ferry ride from Palermo to Naples in 1939.

## Slide 3: Properties of the Neutron
Neutrons are NEUTRAL particles
They are highly penetrating:
can be used as nondestructive probes
can be used to study samples in severe environments
Neutrons have a MAGNETIC moment
They can be used to:
study microscopic magnetic structure
study magnetic fluctuations
develop magnetic materials
Neutrons have SPIN
They can be:
formed into polarized neutron beams
used to study nuclear (atomic) orientation
used for coherent and incoherent scattering.
The ENERGIES of thermal neutrons are similar to the energies of elementary excitations in solids.
The WAVELENGTHS of neutrons are similar to the distance between atoms:
They can determine
atomic structure
structural transitions
Neutrons "see" NUCLEI
They are
sensitive to light atoms,
can exploit isotopic substitution
contrast variation to differentiate
molecular (sub)structures.
*
“highly penetrating”  weakly interacting …
Notes: They are neutral and nuclear (and that light atom sensitivity without the Paraffin Chadwick’s ionization chamber saw nothing we very little –so putting lots of them round a hgih atomic number nucleus electrons mostly don’t care …)

## Slide 4: (Thermal) Neutron Bragg Diffraction
X-10 Graphite reactor 1943
(Clinton Engineering works)
….
E.O. Wollan and C. G. Shull
1945/1946 …
"Scientists at Oak Ridge were very anxious to find real honest-to-goodness scientific uses for the information and technology that had been developed during the war at Oak Ridge and at other places associated with the wartime Manhattan Project.”
Clifford Shull speaking about arriving to work at “Clinton Laboratories” in 1946
An obvious consequence of things already well known in 1932:
X-ray diffraction (1912 the Braggs, father and son)
& de Broglie wavelength hypothesis confirmation (1927 Davisson & Germer)
& Chadwick’s determination of the neutron mass (1.0067 mp)
1936 Mitchel & Powers and Halban & Preiswerk – Bragg diffraction demonstration
Early problems: source & interaction strengths, neutron energies & neutron detection
& WWII …
XRD
from
UC
1948
after n mods
Notes: 1936 – demonstration of quantum wave mechanics …
Weak interaction strength problem as a useful probe ?
Same sort of quasi-problems that Darwin’s formulation of XRD would affect x-ray Bragg, fortunately most crystals are as good as the one in his Darwin’s math – which is why we now have the phrase ideally imperfect …
Wollan was one of Fermi’s team on CP-1 the first “health physicist”, (he invented the pocket sized electroscope ionization chamber dosimeters that are still in use even today )
but may have been outside doing rad survey when the pile went critical and worrying whether his shielding calculations were correct …
Shull live long enough to share the Nobel Prize with Brockhouse (3X) in 1994, which is near a record 1946-1994=48yrs for work to prize (held by Emile Ruska electron microscope 1933-1986=53yrs)
In both cases I’d note that their techniques had enabled several Nobel Prizes well before thay got theirs (Flory – Gaussian polymer chain in melt confirmation definitely an argument, de Geenes also)
Take home lesson: if your ambition is to win a Nobel Prize there are quicker ways of doing it than developing particle scattering or optics techniques …
Also (maybe add later)
1946 - First observation of total external reflection of neutrons by Fermi and Zimm.
1947 - Fermi and Marshall performed detailed studies using monochromatic neutrons to measure scattering lengths of materials => scatt length/x-sections for thermal n (and was Bragging as well)
Neutron Reflection then continued to be used to study bulk scattering lengths.
But Shull and Wollan were the ones who really pioneered ns as a materials science technique Antiferromagnetism – spin matters for neutrons

## Slide 5: What is a neutron scattering instrument ?
Incident Wavevector
Source
Sample
Detector
Scattered Wavevector
ki
kf
Neutron scattering measurements probe nuclear positions and motions in a sample from determinations of the wavevector change (Q) and the energy change (E) of the neutron - S(Q,E).
What do we need to accomplish this?
  1. A source of neutrons
  2. Selection of the wavevector of the incident neutrons (ki)
  3. a) An interesting & otherwise intractable question (n/$/t low)
  3. b) A “large” sample (n interactions are generally “weak”)
  3. c) An otherwise well characterized sample (n/$/t low)
  4. A neutron detector – or lots of them (… n/$/t low)
  5. A method for determining scattered n wavevectors (kf)
Notes: Slecting Neutrons of the right wavelength / energy
Also nice to know as much as you can about your sample beforehand – neutrons are expensive and hard to get to … but needs must

## Slide 6: Wave-particle duality machine
Sample interactions …
wave … mostly
Instrumentally
Paths, flight timing,
collimation, detection …
particle … mostly
de Broglie
Notes: Non-relativistic ….
Mostly for samples … we sometimes have recoil to consider in a way that we wouldn’t do for a ‘normal” wave, but not as it happens unlike Mossbauer …
Mostly for beams, well xtal analyzer interactions, detection is sorta quantum …. For ultracold neutrons diffraction at slits occurs/matters
This can make you feel cool – except for leading to neuty – Gatlinburg conference 1986

## Slide 7: Nuclear Reactors – continuous
Neutrons produced by fission (usually 235U)
Fission requires moderation (~“thermal” neutrons)
Neutrons at a more or less constant rate
Very high neutron fluxes ~1015 n/cm2/s
High Flux
Isotope Reactor
(HFIR – ORNL)
Notes: Neutrons generated in a volume of about 50litre
~10E9 fast n generated per Joule of heat

## Slide 8: Spallation Sources – mostly pulsed
Neutrons produced by nuclear spallation
Accelerator beams usually pulsed => time structure
Neutron energies must be moderated (a lot)
Extremely high peak fluxes ~1016 n/cm2/s
Spallation Neutron
Source
SNS - ORNL
Notes: Neutrons generated in a volume of about 2litre
~10E10 fast n generated per Joule of heat
So actually heat densities aren’t that different per neutron, but actually lower

## Slide 9: Primary characteristics of the sources
SNS
HFIR
Peak flux of the SNS is about 10x the flux of the HFIR
(60 accelerator proton pulses per second)
HFIR flux is about 15x the time averaged flux of the SNS
Time [μs]
Flux n/cm2-s
Notes: Look in the ms around an SNS pulse and the timei integrated neutrons are about the same Note next SNS pulse is nearly 16 graph (15.6) widths away …
After a half a ms or so SNS outputed has dwindled HFIR keeps on chugging

## Slide 10: Neutron Moderation: GeV & MeV  eV & meV
Most HFIR instruments beam tubes
view water cooled Be reflector/moderator
or LH2 Cold Source (T~20K)
SNS Moderators
Neutron Wavelength (Å)
Source Brightness
(n/cm2/sterad/Å/MW/pulse)
Neutrons bounce off something light (better energy transfer) maintained
at a temperature consistent with the energy/wavelength desired
H2O
HFIR Cold Source
lH2 flow head vessel
“Thermal” 2 Å
kBT ~ 25 meV
“Cold” 6 Å
kBT ~ 2 meV
Notes: Reactors have an “advantage” in that although they generate more heat per neutron the nature of sustained fission chain reactions means that the water and reflector design is already trying to cool the neutrons down to thermal energies
NB either way they have to deal with a lot of heat – these are flow through system … in pretty hot environments (radiation heating is a problem – must at least lN2 cool the HFIR cold source while the reactor is running lest its walls mel …)
HFIR moderator is small and thin (constrained to use a small beam tube, volume <1l)
Other reactor moderators tend to use tubs of lD2 ~20-30l circulated through of course
Not coupled moderatior lH2 spectrum lambda peak ~2.5-3A => E~10meV  T~ 290*10/25 ~100K (i.e. not fully moderated – tradeoff between getting them cold and getting them out)
Not really relevant to soft matter, but sometime studies need neutron higher energy neutron than reactor water a hot source 2400C Graphite – there is one at ILL for instance

## Slide 11: Determining the wavevectors: ki & kf
Diffraction and selectors – Reactors mainly
Time-of-flight & Choppers – Pulsed sources mainly
(Focusing)
Cu(200) crystal
monochromator
ILL France
Velocity Selector
Astrium GmbH
Time-of-flight
Time zero from pulse
TOF ~ 10 ms
ΔTOF ~ 100 μs
(moderator dependent)
Leakage time [μs]
Notes: Bragg scattering …
Techniques that rely on the sped of the neutron
Velocity selector operates like a fan – fan analogy - blades here are 10BCarbide to eat neutrons
4A neutron travels at about a km/s … so effective fan speeds a little lower than this
And finally direct time of flight starting from spallation source pulse to when the netron is detected know path length source-sample-detector (or you can do it from a fast chopper obviously)

## Slide 12: So, why neutrons ?
1. Neutrons (now) have the right wavelength
2. Neutrons see the Nuclei
3. Neutrons see light Atoms next to Heavy Ones
4. Neutrons measure the Velocity of Atoms
5. Neutrons penetrate deep into Matter
6. Neutrons see Elementary Magnets

## Slide 13: “Liouville’s theorem” (Gibbs 1902)
http://en.wikipedia.org/wiki/Liouville's_theorem_(Hamiltonian)
“ … the phase space density of non-interacting particles
cannot be increased by conservative forces.”
To increase intensity at a given (Q,E) we must
increase the phase space volume
Simple Instrumental Implication:
It costs flux to increase resolution and it costs resolution to increase flux
~parallel
Small Δk3 high resolution
Need large sample V
converging optics
smaller sample V
Larger Δk3 lower res.
Thermodynamic limit, so in those terms you tend to “lose” more than you “gain”
Drives the multiplicity of neutron instruments
Notes: Obviously the requirements of one measurement large scale structures small angles at reasonable resoultion SANS scattering may be smaller that the angular resolution width of a wide angle diffractometer
Also the range of handlles on neutrons structure & dynamics (Q and E) and different handles Xtal diffraction, TOF, spin precession to characterize and find out what interaction with sample did to an incident neutron

## Slide 14: Steady state use of the neutron spectrum (Elastic scattering example) Continuous source
Sample
Cts
Source
Detectors
Velocity
Selector
Moderated
“White” Beam
Monochromatic
Beam
Q = 4πsinθ / λ
uses narrow spectral band
Scattering
Angle 2θ
Notes: SANS the classic soft matter neutron instrument

## Slide 15: Steady state use of the neutron spectrum (Elastic scattering example) Continuous source
Sample
Cts
Source
Detectors
Velocity
Selector
Moderated
“White” Beam
Monochromatic
Beam
Q = 4πsinθ / λ
uses narrow spectral band
Scattering
Angle 2θ
NG7-SANS NIST Center for Neutron Research
Notes: SANS the classic soft matter neutron instrument
Similar to that 2X widea angle x-ray instrument that Shulla nd Wollan adapted for neutrons at X-10,
But much bigger as interested in larger structures – longer wavelength cold neutron &l a 20m sample to detector distance to measure small scattering angles

## Slide 16: TOF use of the neutron spectrum (Elastic scattering example) Pulsed source
Know Time and Distance
Source-Sample-Detector
Q = 4πsinθ / λ[D,TOF]
uses a wide spectral band
2θ
Sample
Detectors
T=0
Notes: How does a spallation source instrument do that same job …
Source + frame chopper system… So called time of flight gain - caveats

## Slide 17: TOF use of the neutron spectrum (Elastic scattering example) Pulsed source
Know Time and Distance
Source-Sample-Detector
Q = 4πsinθ / λ[D,TOF]
uses a wide spectral band
2θ
Sample
Detectors
T=0
EQ-SANS Spallation Neutron Source
Notes: How does a spallation source instrument do that same job …
Source + frame chopper system… So called time of flight gain - caveats

## Slide 18: TOF use of the neutron spectrum (Elastic scattering example) Use of a wide spectral band allows a wider Q range to be accessed in “one shot”: “Time of Flight Gain” which often makes up for that lower integrated flux in generally wider “survey” of Q (and E for inelastic) Caveats: “Frame overlap” – fast neutrons from later pulses catch slower ones from this solutions: frame choppers (usually a system thereof) and/or some long λ reflection filtering - often also need to shield from unmoderated fast neutron and gamma flash of those succeeding pulses (“T-zero”)
Sample
Detectors
Know Time and Distance
Source-Sample-Detector
Pulsed
Source
T=0
Notes: So called time of flight gain – caveats
Some times you only have one shot. This has been tried for neutron powder scattering from things that are essentially exploding. Not really an option for soft matter – pity.

## Slide 19: Inelastic: Classic Inelastic Reactor Instrument “Triple Axis Spectrometer” (Brockhouse 1955 …)
Source
Sample
Detector
“White” Beam
Monochromatic Beam
Transmitted Beam
Scattered Beam
Scattering
Angle
Bragg
Monochromator
Crystal
Transmitted Beam
ki
kf
Transmitted
Beam
Detector
Analyzer
Crystal
Momentum transfer Q = ki – kf
Energy transfer E = Ei - Ef = (h2/2mn)(ki2– kf2)
Looks at a spot in (Q,E) – but you can scan closely where you want …
kf
Guthoff et al., PRB (1993)
2X Clinton Pile
HB1-3X HFIR ORNL
Notes: Basically you take the two axis design that is pretty much what Shull and Wollan and others were using for elastic scattering structural measurements, add another (monochromator) xtal, et voil – dynmaics!
Gave the definitive answer on the question of whether phonons as quantum excitations of the elastic interaction fields between (in crystals at least) really even existed …
Bertram Brockhouse shared the Nobel Prize with Cliff Shulll in 1994 because he invented this thing … at NRU in 1955/6 at NRX in Canada. Wollan had died some time before unfortunately.

## Slide 20: Inelastic example: Simplified Pulsed Source Chopper Spectrometer
Sample
Source
Detectors
Monochromate
incident beam
(Constant ki)
TOF data reduction
Wide swath of (Q,E)
(data from ARCS-SNS)
Fermi
Chopper
Notes: Fermi chopper is a faster version velocity selector
Helical slots through a vertical rotating drum, but same general principle … sychronize with pulse TOF gives you a ~monochromaticm pulsed beam
The cut you get through qE is the cut you get, whether you want all that info or not – remember we cant tune for everything – can take while, but gentleman’s instrumentation …
This was all rather general – the choppers are used on some instruments at reactors and this example is a so-called direct geometry TOF instrument – meaning that it determines incident wavelengths – in the next lecture by Ken Herwig will talk about an indirect or inverted TOF instrument where final wavelength/energies are determined by crystal diffraction which with TOF information is used to determine the incident emergy … essentially a lot of those interactions of neutrons with matter can be exploited to create instrument components – as my old supervisor used to say “Yesterday’s sensation is tomorrow’s calibration” and you should note a few of these emerging in the next few miniutes …

## Slide 21: Inelastic example: Simplified Pulsed Source Spectrometer Indirect or Inverse Geometry TOF
Sample
Detectors
Detectors
Monochromate scattered beam (Constant kf)
Again TOF data reduction => S(Q,E)
Example: BASIS-SNS (Herwig lecture)
Analyzer
Crystals
kf
T=0
Notes: Fermi chopper is a faster version velocity selector
Helical slots through a vertical rotating drum, but same general principle … sychronize with pulse TOF gives you a ~monochromaticm pulsed beam
The cut you get through qE is the cut you get, whether you want all that info or not – remember we cant tune for everything – can take while, but gentleman’s instrumentation …
This was all rather general – the choppers are used on some instruments at reactors and this example is a so-called direct geometry TOF instrument – meaning that it determines incident wavelengths – in the next lecture by Ken Herwig will talk about an indirect or inverted TOF instrument where final wavelength/energies are determined by crystal diffraction which with TOF information is used to determine the incident emergy … essentially a lot of those interactions of neutrons with matter can be exploited to create instrument components – as my old supervisor used to say “Yesterday’s sensation is tomorrow’s calibration” and you should note a few of these emerging in the next few miniutes …

## Slide 22: Other Choppers: pulse selection, frame & “T-zero”
Path-time graph for butterfly frame chopper choosing second frame (16-32Å) frame SPEAR reflectometer (12.38m, 20Hz)
LANSCE Spallation Source, Los Alamos
(Very Simple) TOF Path-Timing Diagram
Pulse and frame disc
choppers - J-PARC
T-zero choppers – Rotating heavy metal to supresss g and fast neutron “flash” when accelerator proton beam hits spallation target
Distance from source
Notes: Why TOF can get a bit complicated, but frame choice can give you a useful variety of wavelength range choices …

## Slide 23: High pressure gas detectors
~75% of detectors for neutron scattering use 3He at ~3-20 atm (with something easier to ionize added).
“Proportional counter”: energetic particles from n reaction ionize gas, HV (few kV) cascades e charge signal to anode.
Usually cylindrical: collecting at V at
both ends of anode gives position sensitivity, limit resolution ~few mm (charge track cloud size).
These detectors are efficient (usually ~90%+), stable, low noise, and have excellent gamma discrimination, and good timing (~μs) … but they may soon become a thing of the past.
BF3 (Shull and Wollan used it) or 10BF3, timing and resolution problems ...
ARCS Area Array
930 3He LPSD’s
Notes: 3He sources (enriched 235U reactors that bother to collect it) drying up
Other customers (anyone who might want to detect a suspicious source of neutrons …)
SNS has enough stockpiled to make maybe one or two more instruments like ARCS
That’s “linear Position Sensitive Detectors” cf MWPSDs many anode/cathode wires in a 3He volume
Hope if fusion gets going – if anyone has seen the movie “Monn” they may have picked up that the easiest fusion path involves 3He
BF3 cross-section lower so lower efficiency, gas horrible (HF is you break one)
ESS new long pulse spallation source is considering there use – instrument consequences, but they already sacrificed initial time resolution for extra intensity so longer instruments already …

## Slide 24: Neutron scintillator detectors – 1
Typically light is collected from a 6Li doped phosphor screens: ZnS (opaque) or GS20 (clear)
Can use traditional nuclear, HEP coincidence analysis technique: analysis of photomultiplier tube array signals (Anger Camera) 1mm x 1mm resolution & very fast. PM tubes gamma “sensitive”. Installed TOPAZ single xtal.
New technique on two SNS diffraction machines (POWGEN, VULCAN) crossed layers of light shifting optical fibres which collect ZnS blue light convert to green to carry to optical detectors. Resolution ~5mm x 60mm & fast. Efficiencies ~50-60%. Gamma rejection of screens deemed “fair” …
ZnS\6LiF screen
GS20\6LiF laser inspection

## Slide 25: Neutron scintillator detectors – 2
Neutron Radiography & Imaging: CCD Camera views a phosphor screen – again usually 6Li/ZnS. Sub mm resolution, but CCD/high-speed video readout is slower than neutron TOF requires. Gamma sensitive (at right the camera views the ZnS screen through right angle mirror reflection), sensitivity is fairly low, linearity is a problem (cameras are developed to human sensitivity which is logarithmic). Tradeoff between resolution, efficiency, g background and optics.
X-ray image
Neutron Image
N vs X complementarity:
IMAGINE - HFIR
Protein Xtal
High resolution imaging and very high resolution diffraction: BaFBr:Eu2+ “storage phosphor” material for radiation imaging either doped with Gd2O3 or a 6Li or 10B salt. Great resolution (<100μm), very slow readout (by laser scan: Photon Stimulated Luminescence). Dynamic range? And once again g sensitivity is a problem.
M. Fleenor ORNL
N Instrument Group

## Slide 26: Neutron shielding & beam shaping
For a good shielding, aperture, chopper, selector material we seek high cross section neutron nuclear reactions producing something more stoppable or less n detectorable: i.e. softer photons (weak g or x-ray=>Pb, steel, concrete) or charged particles (best)
Some that have been used:
n + 3He → (4He)* → 1H + 3H, Q=0.77MeV (3He~1.4x10-4%) sa= 5.33kb
n + 6Li → (7Li)* → 4He + 3H, Q = 4.8 MeV (6Li~7.5%) sa= 0.94kb
n + 10B → (11B)* → (7Li)* + 4He, Q = 2.3 MeV 0.5MeV g, 94% branch
→ 7Li + 4He, Q = 2.8 MeV, 6% branch (10B~20%) sa= 3.84kb
n + 113Cd → (114Cd)* → 114Cd + g, Q~8MeV (113Cd~12%) sa= 21kb
n + 157Gd → (158Gd)* → 158Gd + g, Q~8MeV (157Gd~16%) sa= 255kb
n + 235U → (236U)* → (fission fragments), Q~200 MeV (235U~0.7%) sa= 0.68kb
Cross sections better for slower/colder neutrons
So mixing in moderator can help, e.g. 6Li-poly,10B-poly, Borax
sa ~ 1/v ~ l ~ 1/T
Also useful: sandwiching of “better” (high sa low g but expensive) in front of
“worse” (lower sa higher g but cheap or structurally convenient) – i.e. 6Li-poly//B-Al//Cd
Notes: NOTE: These are ISOTOPE specific
1/v prop lambda – easier to detect a slower neutron (until it is so slow it can’t get into the detctor)
Charged particles are better because gammas tend to get away from you just like the neutron was going to
Don’t want a high cross-section for other rad esp x gamma – in many cases detector development has faltered because a lot of designers coming from the x-ray world haven’t realized how many photons there are in the world for every neutron and it is almost more important not to detect gammas rather than to detect neutrons
ASIDE Monitors low pressure 3He 235U and N (example N x-section – why neutrons are bad for your health QF~10)

## Slide 27: Incident beam monitoring
When absolute scattering power matters (SANS, NReflectivity …) or relative normalization between scans does (i.e. nearly always)
Low efficiency (<0.1%) transmission detectors in pre-sample path*
n + 3He →(4He)* → 3H + 1H, Q=0.77MeV (3He~1.4x10-4%) sa= 5.33kb
n + 14N → (15N)* → 14C + 1H, Q = 0.63 MeV (14N~99.6%) sa= 1.9 b
(so can just use normal He or N2 in a gas detector …)
n + 235U → (236U)* → (fission fragments), Q~200 MeV (235U~0.7%) sa= 0.68kb
Reactor power stable ~%, but n flux rises as fuel burns. Accelerator operation is somewhat spottier:
He monitor ORDELA Inc.
Fission monitor
* g insensitivity
a must here
SNS P&E vs t
Notes: S can just use normal gas detector design and use normal He or less 3He enriched or N2.
I draw your attention to this second option which is actually a newish thing in monitors, because this reaction is the reason that neutrons radiation gets its 10-20 quality factor from health physicists
Low as it is it is very high for the stuff of which you are made (capture xsec in millib) and N is the N in DNA and _+ you also contain a lot of hydrogenous compounds that moderate neutrons so 1/v is not your friend there …
Flux goes up because 235U is a thermal neutron absorber. As it burns there is less in the way ~5-10% rise at HFIR dependending on what they are irradiating in the core an reflector …
Unlike reactor which are pretty much a few detectors coupled to control rod motors and a tub of water with a pump, accelerators are very complex beasts with thousands of components that have to work
In concert, so there is a lot of tweaking and aiming and the occasional dropout – short drops are the worst – I think the period on the chart is 15m …
Gamma rejection required because these are up front – close to the source and your primary shielding with a pretty good unshielded view

## Slide 28: Neutron transport
Convention for guide mirror angles
m=0.1 deg/Å angle, i.e. relative to θc Ni
Multichannel Curved Guide
Fabricated by Swiss Neutronics
Guide Installation at ISIS UK
The penetrating power of neutrons can make this relatively easy. Thermal neutrons can go a long way in dry air (<10% loss/m for 4Å neutrons). Long flight paths and “get lost” tubes for transmitted beams do require vacuum.
Windows: Aluminium, Quartz, Sapphire & Silicon all have good thermal neutron transmission. Single crystals of the last three can be used to avoid losses due to Bragg scattering. Diamond is also good, and soon affordable?
To transport neutrons a long way from the source (to lower background or just get space) we can use total external reflection in evacuated “neutron guides”:
R~2W/θc[λmin]
2
Notes: Essentially just a reflecting pipe for neutrons – same principle as total internal reflection guiding light in an optical fibre … (external since neutron refractive index is <1)

## Slide 29: Selecting/Analyzing & Manipulating Neutron Spin
Fe
Si
Si
with
B
B
Neutron spin polarizing Supermirrors:
Heusler AlCu2Mn xtal
in B field yoke
Polarizing Bragg reflection
3He Cell
Unpolarized
Neutron Beam
Polarized
Neutron Beam
Xtal Si windowed 3He cell (D17 – ILL)
3He spin polarized cell
Non-adiabatic 90° B change
Neutron Larmor precession
2.91 kHz/G
17 G-cm will “π-flip” a 4Å neutron
passing through the coil
Mezei type Flipper
Polarized neutrons, mirrors, interface reflectivity … (Majkrak and Gaulin)
Notes: Spin polarization can be guided adiabatically – which means that the rotation rate of a magnetic field the neutron sees (remember is is moving through the field) is slow compared to the Larmor precession speed of the magnetic field it is in.
Typically a guide filed will be 20-50G – higher if you want to cope with stray fields from an adjaecnt device - they always say avoid zero fields, but what they really mean is avoid crossing regions where the neutron sees field changes that are confused wrt Larmor frequency …
Mezei flipper is one method - it is wavelength narrow – others Drabkin flipper, RF flipping as per NMR are not so much … usually eat more beamline

## Slide 30: Neutron Spin Optics
A neutron’s magnetic moment and (antiparallel) intrinsic angular momentum undergoes precession around the perpendicular component of an applied magnetic field and will be guided by the parallel component … (some quantum caveats apply)
Precession of polarization is a 3rd way to determine
neutron wavelengths and/or energies
Can be used as a phase “clock” and in very carefully controlled magnetic field conditions this can be used in studies of soft matter dynamics to resolve very small energy changes (meV…neV) – akin to NMR and DLS. This is Neutron Spin Echo.
More recently these techniques have been adapted to resolve very small angular deviations in
the ultra-small angle scattering range (SESANS) and not quite specular reflection (SE Grazing Incidence SANS – SERGIS)
n Gyromagnetic ratio
Phase:
D
D
D
B
B
B
Notes: Neutron Magnetic Resonance, Dynamic Light Scattering ….
As it happens Incoherent scattering (Suny’s lecture) randomizes polarization, coherent scattering maintains the direction – so can in principle be used to get rid of incoherent SANS background.
Has been demonstrated, but incoherent scattering is isotropic – very well behaved and flat – so the costs of polarization (lose at least 50% of neutrons the moment you run your incident beam
through a polarizer), efficiencies of transporting polarization, and the statistical costs of the combining multiple measurements you have to do to

## Slide 31: Sample environment – general considerations
Again the penetrating power of neutrons can make this relatively easy.
Weak scattering power and relatively poor detector resolution means that means instruments tend to be large and have a largish sample position that does not have to be in vacuum. The typical sample area on a modern SANS machine (30m+) allows for about 1m or so of sample environment beam path. Large installations and ancillary access possible.
2.5 cm
Reflection geometry cell for NR and SANS studies of adsorbed layers and colloidal solution morphology near surfaces. Incident and scattered beams pass through 8cm of crystal quartz slab. Transmission at 4Å ~75%
Windows and sample environment enclosures: Aluminium, Quartz, Sapphire & Silicon all have good thermal neutron transmission. Single crystals of the last three can be used to avoid losses due to Bragg scattering – if necessary you can send a beam through many cm.
kfR
ki
QR
Quokka SANS sample position – in-situ SANS Rheometer
- OPAL ANSTO
Quokka SANS sample position – 11T Horizontal field magnet
- OPAL ANSTO
Quokka SANS sample position – in-situ SANS Rheometer
- OPAL ANSTO
Notes: It may already be available – consult the web pages of the neutron scattering facilities – often if you seem interesting enough theIR will work with you to develop SE for routine use

## Slide 32: So, what is a neutron scattering instrument ?
Incident Wavevector
Source
Detector
Scattered Wavevector
ki
kf
ki
Wavelength range
- moderator
ambient, cold, hot
Monochromatic
“white” – moderated
“pink” – filtered moderator
Pulsed
– wavelength frame ?
Polarized ?
Sample
Alignment matters:
- Single xtal, flat interface
Or doesn’t:
Powder, amorphous solid,
liquid, solution, gas
Length (k) or energy (E) scale of interest ?
Scattering power
& transmission
Sample environment:
T, P, B, …
Phase changes: scale, rate, reversibility …
kf
Structure, kinematics or dynamics … ?
Single wavelength or range ? Linear or area detector ?
Resolve or filter its energy?
- Xtal or TOF chopper?
Elastic, quasi-elastic
or inelastic. Range?
High resolution, low resolution, counting time ?
 Sample size, stability, scattering power
Analyze final polarization?
Notes: Selecting Neutrons of the right wavelength / energy
Also nice to know as much as you can about your sample beforehand – neutrons are expensive and hard to get to … but needs must

## Slide 33: Elastic Neutron Scattering Instruments
Small Proteins
Micelles
Colloids Vesicles
Critical Phenomena
Membranes
Liquids and Glasses
Crystal (inc protein),
Fiber,
and
Magnetic
Structures
Metals Alloys Zeolites
Magnetic Films
and Domains
Biological Structure
Viruses
Polymers
SANS
REFLECTOMETERS
DIFFRACTOMETERS
  Techniques:
  Powder diffraction
  Single Crystal diffraction
  Small Angle Scattering
  Reflectometry
Based on diagram by Roger Pynn
d (Å)
Longer Length Scales
Q(Å-1)
  To determine the
  average structure of materials.
  How the atoms are arranged or are rearranging.
  Structure & Kinematics
Notes: Reflectometers are often polarized, as are some SANS

## Slide 34: Inelastic Neutron Scattering Instruments
0.1
1
10
100
d(Å)
Neutron Induced
Excitations
Hydrogen
Modes
Molecular
Vibrations
Intenerate
Magnets
Crystal
Fields
Spin
Waves
Lattice
Vibrations
Coherent
Modes
Electron-Phonon
Interactions
Molecular
Motions
Aggregate Motion
Polymers and Biological Molecules
Critical Scattering
Diffusion
Molecular
Reorientation
Based on diagram by Rex Hjelm
Slower
Motions
Longer Length Scales
Spectrometry Techniques
Direct Geometry TOF
Triple-Axis
Indirect Geometry TOF
Backscattering …
Neutron Spin-Echo
Atomic, molecular or assemblage dynamics:
phonons, magnons, diffusion …
How are the atoms are moving within a given structure or morphology.
Dynamics
THz
GHz
MHz
Notes: Over ~8 orders of magnitude in energy/time and length scale

## Slide 35: Note: Big facilities, long user lead times …
HFIR – ORNL SNS - ORNL
What 1 or 2 billion $ buys in terms of neutron sources:
In pure flux terms neither compares very well with a the photon brightness of a bench top x-ray source. (Let alone a rotating anode or synchrotron) & neutron interactions tend to be fairly weak …
Notes: SNS is now about 5 years old, HFIR about 45
SNS cost about 1.5 billion dollars, supposing you could build it again HFIR would cost about the same, and yet … we do smallish science on them

## Slide 36: Where can I do neutron scattering ?
Adapted from www.veqter.co.uk website
Major scientific neutron scattering sources worldwide
Notes: At the moment about 2:1 reactors over accelerators (various) – reactors are aging (ANS was supposed to be built in the 90’s) new sources ESS in Lund Sweden circa 2020+ … and Los Alamos may soon disappear
You can do neutron scattering all over the world, but this is apretty complete list of all neutron scattering sources – neutron radiography aside – which tend by their very nature to be major when compared with other instrument suites

## Slide 37: The neutron zoo in information and real space
Covers 7-8 orders of magnitude in length and energy (time), 2-3/instrument
HFIR and SNS instruments.
Similar extensive suites at NCNR-NIST, ILL (France), ISIS (UK), LANSCE-LANL, …
Notes: Why so many types of instruments & why are neutron beam rooms and guide halls so crowded ?
Types of sources (and their numerical paucity)
Xtal, TOF, spin precession beam hanndling
Sample environment
Q,E range for sample – the eternal fight againt Liouville’ theorem
Efficient use of instrument optimized over narrowish Q,E,ranges

## Slide 38: Conclusion: Why neutrons ?
“ … because [neutron scattering] provides information about the structure of materials that cannot be obtained in simpler, less expensive ways.”
1. Neutrons have the right wavelength
2. Neutrons see the Nuclei
3. Neutrons see light Atoms next to Heavy Ones
4. Neutrons measure the Velocity of Atoms
5. Neutrons penetrate deep into Matter
6. Neutrons see Elementary Magnets
7. And (because) scattering is inherently statistical
“ … in spite of all these penalties of a signal-limited technique, neutron scattering continues to occupy an important place among the panoply of tools available to study materials structure because it can often provide information that cannot be [simply] obtained in any other way.”
R. Pynn “Neutron Scattering—A Non-destructive Microscope for Seeing Inside Matter”
in Neutron Applications in Earth, Energy and Environmental Sciences, L. Liang et al. (eds.),
Neutron Scattering Applications and Techniques, Springer 2009
Notes: A picture is worth a thousand words … a Scattering pattern is worth a thousand pictures …
Insertion & italics are mine My italics …
The effect you seek is hidden behind another (similar) system, buried under a layer that masks other probes, or is simply deep inside whatever sample environment equipment
Neutron scattering can often let you peer through that fog of other stuff (that’s what highly penetrating means) and has advantages & methods that can allow the signal from the
system that you are interested in stand out clear.
Next lecture SANS, and I’ll try to put a little more flesh on the bones of those assertions … and now for whatever time we have left I’ll try to answer questions (don’t say that don’t steal too much thunder from the next presentation)

## Slide 39: Acknowledgements
Material has been borrowed and adapted from presentations and publications by:
Lee Robertson, Steve Nagler, Ken Herwig, Lowell Crow, Rick Reidel, John Katsaras, Ian Anderson, Tom Mason, Mike Fleenor …
Roger Pynn, Jack Carpenter, Paul Butler, Boualem Hammouda, Bruce Gaulin, Jim Rhyne, Ross Stewart, Rob Dalgleish, Ken Andersen, Rex Hjelm …
& whoever they borrowed and adapted from (in some cases that was me).
General References:
“Elementary Scattering Theory: For X-ray and Neutron Users”, D.S. Sivia, Oxford (2011)*
“Introduction to the Theory of Thermal Neutron Scattering”, G.L. Squires, Cambridge (2012)
“Neutron, X-Ray and Light Scattering: Introduction to and Investigative Tool for Colloidal and Polymeric Systems” (1991)
“Neutrons, X-Rays and Light: Scattering Methods Applied to Soft Condensed Matter” (2002)
P. Lindner and T. Zemb, Elsevier North-Holland
“Neutrons in Soft Matter”, T. Imae, T. Kanaya, M. Furusaka and N.Torikai, Wiley (2011)
“Dynamics of Soft Matter: Neutron Applications”, Springer (2012)
V. García Sakai, C. Alba-Simionesco and S.-H. Chen
“Practical Neutron Scattering at a Steady State Source”,
T. Heitmann & W. Montfrooij, Mizzou Media (2012)*
http://www.ncnr.nist.gov/summerschool/ss13/pdf/NeutronScatteringPrimer.pdf R.Pynn (1990)*
“Handbook of Neutron Optics”, M. Utsoro and V.K. Ignatovich, Wiley (2010)
*most “accessible”
Notes: Sometimes that was me – bringing a certain closure to the process

## Slide 40: Neutron Scattering Source Websites
http://neutrons.ornl.gov
http://www.ncnr.nist.gov
http://lansce.lanl.gov
http://www.murr.missouri.edu/rd_material_sciences.php
http://www.indiana.edu/~lens/
http://www.ill.eu
http://www.isis.stfc.ac.uk
http://www.frm2.tum.de
http://www.psi.ch/sinq/
http://www.tnw.tudelft.nl/en/cooperation/facilities/reactor-instituut-delft/
http://www-llb.cea.fr/en/
http://www.ansto.gov.au/ResearchHub/Bragg/index.htm
http://flnp.jinr.ru/562/
http://www.nti.org/facilities/9/
… a more exhaustive list: http://www.neutron.anl.gov/facilities.html
Notes: Sorry Asia. Rather a lot of small reactors do radiography and or some kind of materials analysis irradiation …